# Zeiss ORION NanoFab (HIM) SOP

Prepared by Xu Yue, Ashley

Before Measurement

- Cool down body/clothing temperature before entering

- *Three windows should be open when beginning the session
-System Manager (click start, wait 1 min)
-Zen (only when "server has started")
-Tools (vacuum screen)

- Check for error messages -> call administrator if there are issues

- Check and fill up logbook -> check trimer

- Check gun temperature < ~ 80K (if greater than 84K, don't use)

- Check liquid N2 remaining time, don’t proceed if <2h

- Check vacuum status (GFIS~E-10, Chamber ~ E-8), 2E-7

- Load sample

  - Sample management

  - "Exchange" - also neutralizes stage

  - "Transfer" to open load lock / chamber wall

  - Insert the fetching rod steadily to couple with the sample holder. Screw the rod till the couple is locked.

  - Pull the sample stage out. When the transfer rod is retracted completely, the "Rod Retracted" indicator on the sample chamber will glow.

  - "Store" to close chamber wall and keep load lock in vacuum <- keep it in "Store" position

  - "Vent"

  - Load samples. Please take photos to mark position, top icon

  - Back to "store" When loading, "Rod Retracted" must be green -interlock

  - wait for the sample chamber to pump to 2E-2 after store, then open the gate door (transfer)

- In stage position tab, ensure X and Y directions are unticked. Untick relative move. Move Z to 0. Then tick Z to 3.65(3.76 calibrated Feb2021) and move to (generally used for 500um thick substrate. always less than 4mm)

- Power settings -> turn on DT detector (recommended to not turn on the gun from here, but from GFIS column control).

- Expand the GFIS column tab.

  - Always check the suitable aperture is inserted before turning on the gun! He apertures are made of Au

  - Start He gas flow

  - *Never use the Extract Voltage

  - Turn on Lens 2 (power supply enable) and slowly increase voltage to ~19kV by tuning the focus knob.

- Check trimer alignment

  - move the stage to standard Au sample (saved as mark22 on site list)

  - rough focus on Au surface: increase magnification to ~95um FOV, tune focus to WD is ~ 9.5mm

  - set scanning parameter (dwell time 0.1us, average "none")

  - go to SFIM mode (bottom on keyboard)

  - Grab a scanning image. Check if the center of the cursor is aligned with designated trimer.

  - switch to GUN TILT mode to align the trimer if needed

  - * Expose the trimer as shortly as possible *

  - Go back to GUN SHIFT mode and normal imaging mode

Parameter settings

    - Open GFIS panel

    - Accelerator voltage default value is 30kV. Decrease 3kV at a time. When changing the accelerator voltage, lens1 and lens2 parameters also will change. If not, press twice SFIM/normal button to reset.

    - Check trimer again. Tune the gun shift accordingly to adjust trimer position. If gun shift value (bottom of GFIS panel) >200urad, reset gun shift to 0 then manually align the trimer by tuning knobs on the column.

    - Beam current should be a little higher than designed because the beam intensity will drop during the experiment. Tune the current spot to achieve desired beam current. Smaller spot => higher current

    - Save probe setting for reference. * probe setting also includes accelerator voltage, so tune down to desired voltage before going to saved probe settings. Name the probe setting as name_material_aperture_voltage

For image acquisition

- Use standard sample to tune focus so we minimize the exposure of samples.

- Go to region of interest (sample map or site list) and start scanning by choose "Continuous". Adjust contrast and brightness as needed (~50%, no auto adjust for Zeiss)

- At sample, change parameters to desired:

  - change spot size till current ~ 1pA (grab in normal imaging mode to reset the current back) How to decide what spot size to use?

  - 2us dwell time

  - turn off average mode

  - scan size 512x512

- Alignment at flat region - do it at the beam current you're interested in

  - Correct astig with knobs

  - Correct optical path with GUN SHIFT knobs and wobbler (change dwell time to 0.1us when open wobbler)

  - tune focus knob

  - Repeat 1-3 to get best resolution

  - Take notes of stig and lens 2 values for reference.

- Focus at the edge first then go to features near the flake of interest. Focus again.

- Stop scanning and go to your sample. Start "Continuous" scanning to find location of interest. Save locations using left panel "+" (Never clear position 5!!)

- Capturing image

  - Line averaging 8/16

  - Dwell time = 2us, if high quality image is needed, dwell time can be increased to 5us.

  - Electron flood: time~600us and adjust crosshair. Flood gun never use On mode, use line

  - Choose "Grab" for image capture

For patterning

- choose patterning tab on top

- focus before patterning to get a sharp edge

- Dose use Ions/cm2

- Frame averaging for Ne patterning

- Click "Expected beam current" - 2.6 pA for polishing

- Ne- 10usec, He 100usec

- Pixel spacing 1nm. Dwell time increase → repeats decrease

- Other patterning parameters?

- For more complicated patterns, use nanofab software (right PC). Check server status on the right PC. It should be connected to the HIM. If not, open communication software NPVE on left PC to reconnect.

- Draw patterns on nanofab software. Adjust parameters. Select all patterns. Press start from the nanofab GUI.

- Nanofab accepts bmp files.

Finish measurement

- Power down detector (watch vacuum change from E-6 to E-10)

- In GFIS column tab, close He gas.

- Slowly tune down Lens 2 voltage to 0 then untick "power enable"

- Wait for the GFIS/Chamber valve to close. Check in the vacuum screen.

- Take out your sample then load the standard sample back to the main chamber. Always keep the sample stage in the main chamber.

Others

- Use built-in plasma cleaner to remove hydrocarbon. How and when?

- Total tilt is between 0-54

- Probe voltage = 4.5/3.8V per Angstrom